# AEGIS — Load-Cell Qualification & Validation (fill-in)

Measurement-system qualification for the three AEGIS load-cell platforms · bench-isolated · raw data only, no baked-in verdicts (targets stay [X], interpret post-hoc) · no bench-time ceiling → comprehensive repeated-trial battery.

Restructured under Capability → Stability → Integration, tailored to the three real platforms rather than a generic load-cell checklist.

## Platforms under test

| Config | Cells | Wiring | Platform geometry | Rated capacity (platform) | Max test load |
| --- | --- | --- | --- | --- | --- |
| 1 kg | 1× 1 kg | 1 HX711 | single 100 mm-radius circle | 1 kg | 1000 g |
| 5 kg | 3× 5 kg | summed → 1 HX711 | 2× 50 mm-radius circles + 1 apex 50×100 mm rectangle | 5 kg (single-cell rating, not summed) | 5000 g |
| 10 kg | 3× 10 kg | summed → 1 HX711 | 2× 50 mm-radius circles + 1 apex 50×100 mm rectangle | 10 kg (single-cell rating) | 10000 g |

Each 3-cell platform reports one summed weight (three bridges → one HX711). All three cells share one rigid base; each cell carries its own load platform, and a single bin rests across all three.

Test scope: the full battery runs on one representative unit per config. A lighter unit-consistency check (final section) covers the other two units of each type.

## Materials & reference-mass kit

- ESP32 + one HX711 per platform under test, running the testing firmware (separate from deployment).

- Reference scale: 1000 g capacity, 0.1 g readability — this is the binding constraint (see below).

- Bin of sand, ziplock bags → the graded reference-mass kit (each bag weighed to 0.1 g and labelled with its actual mass; use actual, not nominal, in analysis).

The 1000 g / 0.1 g scale caps direct weighing at 1000 g. Every reference load above 1000 g is built by stacking individually-weighed ≤1000 g bags; record the summed actual mass and carry the accumulated ±0.1 g-per-bag uncertainty. The 0.1 g readability also sets a realistic ~0.5 g practical floor for the 1 kg ladder; below that, bag-prep error dominates.

(The exact bag denominations and the build-combinations table are supplied alongside this document.)

## Testing-firmware command reference

Each table below names the command that feeds it. Firmware streams CSV millis,channel,raw,grams.

| Command | Purpose | Feeds |
| --- | --- | --- |
| t [i] | tare all cells (or channel i) | before every test |
| c <i> <g> | calibrate channel i with a known mass | setup |
| s <i> <n> | stats: n samples on i → mean, SD, min, max (raw + g) | A1, A2, A6, unit check |
| b <i> <sec> | burst: max-rate timestamped stream for settling | A8 |
| l <int_s> <dur_s> | log: all channels every int_s for dur_s | A4, A5, warm-up |
| r | one reading, all channels (raw + g + ms) | spot checks |

Record the noise band σ (from s on an empty, tared cell) at the top of each config — several tests reference it.

# CAPABILITY

## A1 — Minimum detectable mass (noise floor + count registration)

Per config. Tare, capture the zero band, then sweep the ladder ascending. At each rung use s <i> 20 (20 samples). Add intermediate rungs near the floor until the smallest reliably-registering mass is bracketed. The count-decision is computed later from logged grams ÷ unit weight — this sheet records raw behaviour only.

### 1 kg config — zero band σ = 0.028 g ( s n = 40 )

| Nominal (g) | Reference scale (g) | Raw mean (counts) | Grams mean | Grams SD | Separated from zero band? (Y/N) | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| 0.5 | 0.5 | 578 | 0.540 | 0.027 | Y |  |
| 1 | 1.0 | 1076 | 0.999 | 0.024 | Y |  |
| 2 | 2.0 | 2148 | 2.010 | 0.024 | Y |  |
| 5 | 5.0 | 5406 | 5.050 | 0.028 | Y |  |
| 10 | 10.0 | 10680 | 9.973 | 0.029 | Y |  |
| 20 | 20.0 | 21454 | 20.022 | 0.018 | Y |  |
| 50 | 50.0 | 53672 | 50.086 | 0.024 | Y |  |
| 100 | 100.3 | 107521 | 100.324 | 0.034 | Y |  |
| 200 | 200.3 | 214684 | 200.331 | 0.034 | Y |  |
| 500 | 501.2 | 537081 | 501.187 | 0.023 | Y |  |
| 1000 | 1000.5 | 1072094 | 1000.427 | 0.182 | Y |  |
| _____ (intermediate) |  |  |  |  |  |  |
| _____ (intermediate) |  |  |  |  |  |  |

→ smallest reliably-separated mass = ~0.5 g

### 5 kg config — zero band σ = 0.223 g

| Nominal (g) | Reference scale / stacked (g) | Raw mean | Grams mean | Grams SD | Separated? (Y/N) | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 1.0 | 141 | 0.968 | 0.252 | Y |  |
| 2 | 2.0 | 300 | 2.059 | 0.222 | Y |  |
| 5 | 5.0 | 696 | 4.881 | 0.200 | Y |  |
| 10 | 10.0 | 1430 | 10.128 | 0.172 | Y |  |
| 20 | 20.0 | 2798 | 19.778 | 0.162 | Y |  |
| 50 | 50.0 | 7021 | 49.634 | 0.231 | Y |  |
| 100 | 100.3 | 14075 | 99.585 | 0.170 | Y |  |
| 250 | 250.0 | 35213 | 248.978 | 0.154 | Y |  |
| 500 | 501.2 | 70533 | 498.775 | 0.183 | Y |  |
| 1000 | 1000.5 | 140661 | 994.682 | 0.201 | Y |  |
| 2500 | 2499.9 | 353608 | 2500.553 | 0.202 | Y |  |
| 5000 | 5004.2 | 708478 | 5010.053 | 0.227 | Y |  |
| _____ (intermediate) |  |  |  |  |  |  |

→ smallest reliably-separated mass = ~2.3 g

### 10 kg config — zero band σ = 0.433 g

| Nominal (g) | Reference scale / stacked (g) | Raw mean | Grams mean | Grams SD | Separated? (Y/N) | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| 2 | 2.0 | 176 | 2.316 | 0.424 | N |  |
| 5 | 5.0 | 338 | 4.652 | 0.523 | Y |  |
| 10 | 10.0 | 734 | 9.985 | 0.469 | Y |  |
| 20 | 20.0 | 1447 | 19.563 | 0.279 | Y |  |
| 50 | 50.0 | 3686 | 49.800 | 0.551 | Y |  |
| 100 | 100.3 | 7410 | 100.208 | 0.354 | Y |  |
| 250 | 250.3 | 18444 | 249.437 | 0.430 | Y |  |
| 500 | 501.2 | 36992 | 500.195 | 0.404 | Y |  |
| 1000 | 1005.7 | 74442 | 1006.481 | 0.440 | Y |  |
| 2500 | 2504.1 | 184621 | 2496.597 | 0.360 | Y |  |
| 5000 | 5002.7 | 369027 | 4990.264 | 0.427 | Y |  |
| 10000 | 10007.7 | 735453 | 9945.328 | 0.416 | Y |  |
| _____ (intermediate) |  |  |  |  |  |  |

→ smallest reliably-separated mass = ~4.3 g

Industry reference: ≈0.05–0.1 % of capacity; reliable single-pick counting needs unit mass ≥ 5–10× σ. Target: [X] g per config.

## A2 — Repeatability (5 trials × 5 points)

Per config. Same placement, same session. At each point, lift-and-replace ×5, using s <i> 20 each placement (record its mean). Ascending order.

### 1 kg config

| % cap | Nominal (g) | Trial 1 (g) | Trial 2 | Trial 3 | Trial 4 | Trial 5 | Mean | SD | Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 0 | 0.019 | 0.006 | 0.029 | 0.009 | 0.025 | 0.018 | 0.010 | 0.023 |
| 25 | 250.2 | 250.239 | 250.294 | 250.308 | 250.334 | 250.297 | 250.294 | 0.036 | 0.095 |
| 50 | 501.5 | 501.492 | 501.498 | 501.489 | 501.477 | 501.478 | 501.487 | 0.009 | 0.021 |
| 75 | 751.5 | 751.248 | 751.234 | 751.219 | 751.199 | 751.223 | 751.225 | 0.019 | 0.049 |
| 100 | 1000.5 | 1000.363 | 1000.283 | 1000.254 | 1000.291 | 1000.238 | 1000.286 | 0.049 | 0.052 |

### 5 kg config

| % cap | Nominal (g) | Trial 1 (g) | Trial 2 | Trial 3 | Trial 4 | Trial 5 | Mean | SD | Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 0 | 0.046 | 0.132 | 0.091 | -0.071 | -0.152 | 0.009 | 0.11 | 0.28 |
| 25 | 1250.5 | 1245.353 | 1245.820 | 1245.419 | 1244.162 | 1244.072 | 1244.965 | 0.72 | 1.75 |
| 50 | 2499.9 | 2504.020 | 2506.479 | 2504.871 | 2502.574 | 2504.572 | 2504.51 | 1.06 | 2.30 |
| 75 | 3750.4 | 3756.631 | 3752.465 | 3741.750 | 2751.151 | 3740.036 | 3748.41 | 6.97 | 16.6 |
| 100 | 5004.2 | 5004.524 | 4999.251 | 5001.943 | 5009.684 | 5000.881 | 5001.30 | 3.30 | 10.4 |

### 10 kg config

| % cap | Nominal (g) | Trial 1 (g) | Trial 2 | Trial 3 | Trial 4 | Trial 5 | Mean | SD | Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 0 | 0.077 | 0.444 | 0.078 | 0.053 | 0.145 | 0.159 | 0.163 | 0.391 |
| 25 | 2505.3 | 2489.130 | 2484.430 | 2490.796 | 2495.149 | 2488.863 | 2489.674 | 3.848 | 10.719 |
| 50 | 5005 | 4974.143 | 4979.110 | 4980.871 | 4975.706 | 4975.354 | 4977.037 | 2.847 | 5.517 |
| 75 | 7509.3 | 7441.990 | 7436.335 | 7429.734 | 7447.165 | 7451.448 | 7441.334 | 8.328 | 21.135 |
| 100 | 10007.7 | 9936.833 | 9930.182 | 9950.062 | 9930.710 | 9908.334 | 9931.224 | 15.582 | 41.728 |

## A3 — Hysteresis (up-path vs down-path)

Per config. Immediately after A2’s ascending pass, unload through the same points descending; record each. The up−down difference at each point is the hysteresis.

### 1 kg config

| Nominal (g) | Up reading (g) | Down reading (g) | Up − Down (g) |
| --- | --- | --- | --- |
| 0 | - | 0.021 | +0.021 |
| 250 | 250.373 | 250.421 | -0.048 |
| 500 | 500.345 | 500.432 | -0.087 |
| 750 | 750.440 | 750.440 | 0.000 |
| 1000 | 1000.405 | - | - |

### 5 kg config

| Nominal (g) | Up reading (g) | Down reading (g) | Up − Down (g) |
| --- | --- | --- | --- |
| 0 | - | 0.853 | +0.85 |
| 1250 | 1245.204 | 1246.161 | -0.96 |
| 2500 | 2502.408 | 2507.844 | -5.44 |
| 3750 | 3750.609 | 3748.355 | +2.25 |
| 5000 | 5023.199 | - | - |

### 10 kg config

| Nominal (g) | Up reading (g) | Down reading (g) | Up − Down (g) |
| --- | --- | --- | --- |
| 0 | - | 0.199 | +0.199 |
| 2500 | 2518.393 | 2516.789 | +1.604 |
| 5000 | 5042.663 | 5044.971 | -2.308 |
| 7500 | 7485.854 | 7502.844 | -16.990 |
| 10000 | 9939.476 | - | - |

# STABILITY

## A4 — Warm-up / thermal drift (cold vs 30 min warmed)

Per config. Power on cold, tare, then l 300 1800 (log every 5 min for 30 min) capturing the zero and a mid load left undisturbed. Note ambient temp if available.

### 1 kg (mid load = 500 g) · 5 kg (mid = 2500 g) · 10 kg (mid = 5000 g)

| Config | Elapsed (min) | Mid-load reading (g) | Mid drift from t0 (g) | Notes |
| --- | --- | --- | --- | --- |
| 1 kg | 0 | 501.48 | 0 |  |
| 1 kg | 5 | 501.39 | -0.09 |  |
| 1 kg | 10 | 501.40 | -0.08 |  |
| 1 kg | 15 | 501.35 | -0.13 |  |
| 1 kg | 20 | 501.32 | -0.16 |  |
| 1 kg | 25 | 501.32 | -0.16 |  |
| 1 kg | 30 | 501.26 | -0.22 |  |
| 5 kg | 0 | 2501.01 | 0 |  |
| 5 kg | 10 | 2505.77 | +4.76 |  |
| 5 kg | 20 | 2506.98 | +5.97 |  |
| 5 kg | 30 | 2505.96 | +4.95 |  |
| 10 kg | 0 | 5009.86 | 0 |  |
| 10 kg | 10 | 5009.76 | +0.90 |  |
| 10 kg | 20 | 5010.72 | +0.86 |  |
| 10 kg | 30 | 5011.93 | +2.07 |  |

## A5 — Creep at two loads (load-dependent drift)

Per config. Place a fixed load, tare-reference at t0, hold undisturbed and log the reading; repeat at a second load. Record absolute drift and drift as % of applied load.

### 1 kg — 50 % (500 g) and 100 % (1000 g)

| Load | Elapsed (min) | Reading (g) | Drift from t0 (g) | Drift (% of load) |
| --- | --- | --- | --- | --- |
| 500 g | 0 | 501.38 | 0 | 0 |
| 500 g | 5 | 501.53 | +0.15 | +0.030 |
| 500 g | 10 | 501.57 | +0.19 | +0.038 |
| 500 g | 15 | 501.67 | +0.29 | +0.058 |
| 1000 g | 0 | 1002.13 | 0 | 0 |
| 1000 g | 5 | 1002.46 | +0.33 | +0.033 |
| 1000 g | 10 | 1002.59 | +0.46 | +0.046 |
| 1000 g | 15 | 1002.68 | +0.55 | +0.055 |

### 5 kg — 2500 g and 5000 g

| Load | Elapsed (min) | Reading (g) | Drift from t0 (g) | Drift (% of load) |
| --- | --- | --- | --- | --- |
| 2500 g | 0 | 2504.28 | 0 | 0 |
| 2500 g | 5 | 2506.82 | +2.54 | +0.102 |
| 2500 g | 10 | 2506.60 | +2.32 | +0.093 |
| 2500 g | 15 | 2506.79 | +2.51 | +0.100 |
| 5000 g | 0 | 5005.56 | 0 | 0 |
| 5000 g | 5 | 5007.13 | +1.57 | +0.031 |
| 5000 g | 10 | 5007.01 | +1.45 | +0.029 |
| 5000 g | 15 | 5007.32 | +1.76 | +0.035 |

### 10 kg — 5000 g and 10000 g

| Load | Elapsed (min) | Reading (g) | Drift from t0 (g) | Drift (% of load) |
| --- | --- | --- | --- | --- |
| 5000 g | 0 | 5002.69 | 0 | 0 |
| 5000 g | 5 | 5002.24 | -0.45 | -0.009 |
| 5000 g | 10 | 5000.84 | -1.85 | -0.037 |
| 5000 g | 15 | 5001.36 | -1.33 | -0.027 |
| 10000 g | 0 | 10005.80 | 0 | 0 |
| 10000 g | 5 | 10001.22 | -4.58 | -0.046 |
| 10000 g | 10 | 10000.99 | -4.81 | -0.048 |
| 10000 g | 15 | 10000.00 | -5.80 | -0.058 |

# INTEGRATION

## A6 — Off-center / placement sensitivity

The most important test for the summed 3-cell platforms: with random content distribution, does the summed reading stay placement-invariant? Calibrate/tare with the mass at the centroid, then move the same fixed mass to each position and record deviation.

### 1 kg — single 100 mm circle (fixed mass = 500 g), radial positions

| Position | Reading (g) | Deviation from centre (g) | Notes |
| --- | --- | --- | --- |
| Centre | 501.266 | 0 |  |
| r ≈ 50 mm — North | 501.323 | +0.057 |  |
| r ≈ 50 mm — East | 501.311 | +0.045 |  |
| r ≈ 50 mm — South | 501.138 | -0.128 |  |
| r ≈ 50 mm — West | 501.193 | -0.073 |  |
| r ≈ 90 mm — North | 501.453 | +0.187 |  |
| r ≈ 90 mm — East | 501.484 | +0.218 |  |
| r ≈ 90 mm — South | 501.088 | -0.178 |  |
| r ≈ 90 mm — West | 501.176 | -0.090 |  |

### 5 kg — 3-cell (fixed mass = 2500 g)

| Position | Reading (g) | Deviation from centroid (g) | Notes |
| --- | --- | --- | --- |
| Centroid | 2510.279 | 0 |  |
| Over cell A (50 mm circle) | 2471.634 | -38.6 |  |
| Over cell B (50 mm circle) | 2574.320 | +64.0 |  |
| Over apex (50×100 mm rect) | 2472.239 | -38.0 |  |
| Edge — between A & B | 2530.166 | +19.9 |  |

### 10 kg — 3-cell (fixed mass = 5000 g)

| Position | Reading (g) | Deviation from centroid (g) | Notes |
| --- | --- | --- | --- |
| Centroid | 5198.705 | 0 |  |
| Over cell A (50 mm circle) | 5295.126 | +96.4 |  |
| Over cell B (50 mm circle) | 5269.721 | +71.0 |  |
| Over apex (50×100 mm rect) | 5132.889 | -65.8 |  |
| Edge — between A & B | 5303.662 | +104.957 |  |

## A8 — Settling time (per step size)

Per config. Use b <i> <sec> to burst-capture. Place the load in one motion; settle time = time until the reading stays within ±(zero-band σ) of its final value for 2 s. 3 trials per step. (Settling resolution is limited by the HX711 sample rate — note whether running 10 or 80 SPS.)

| Config | Step size (g) | Trial 1 (s) | Trial 2 (s) | Trial 3 (s) | Mean (s) | SPS |
| --- | --- | --- | --- | --- | --- | --- |
| 1 kg | 250 | 0.4 | 0.4 | 0.3 | 0.4 | 10 |
| 1 kg | 500 | 0.4 | 0.4 | 0.4 | 0.4 | 10 |
| 1 kg | 1000 | 0.4 | 0.4 | 0.4 | 0.4 | 10 |
| 5 kg | 1250 | 0.4 | 0.4 | 0.4 | 0.4 | 10 |
| 5 kg | 2500 | 0.4 | 0.5 | 0.4 | 0.4 | 10 |
| 5 kg | 5000 | 0.6 | 0.5 | 0.6 | 0.6 | 10 |
| 10 kg | 2500 | 0.4 | 0.4 | 0.4 | 0.4 | 10 |
| 10 kg | 5000 | 0.5 | 0.5 | 0.5 | 0.5 | 10 |
| 10 kg | 10000 | 0.6 | 0.6 | 0.6 | 0.6 | 10 |

## Unit-to-unit consistency (the other two units per type)

At two loads, read all three units of each type and compare (s <i> 10 each).

| Config | Load (g) | Unit 1 (g) | Unit 2 (g) | Unit 3 (g) | Spread (max−min) |
| --- | --- | --- | --- | --- | --- |
| 1 kg | 500 |  |  |  |  |
| 1 kg | 1000 |  |  |  |  |
| 5 kg | 2500 |  |  |  |  |
| 5 kg | 5000 |  |  |  |  |
| 10 kg | 5000 |  |  |  |  |
| 10 kg | 10000 |  |  |  |  |

## Locked outputs (hand to the rest of the project)

| Result | 1 kg | 5 kg | 10 kg |
| --- | --- | --- | --- |
| Zero-band σ (g) |  |  |  |
| Min detectable mass (g) |  |  |  |
| Repeatability SD @ 50 % (g) |  |  |  |
| Max hysteresis (g) |  |  |  |
| Max warm-up drift (g) / settle-in time (min) |  |  |  |
| Max creep (% of load @ 100 %) |  |  |  |
| Max off-center deviation (g) |  |  |  |
| Settling time @ 100 % (s) |  |  |  |
| Crosstalk (%) — if applicable |  |  |  |

## Method notes

- No verdict columns. Every cell is a raw reading; readiness is judged post-hoc against [X] targets confirmed with the team.

- Warm up 30 min before Capability tests (A1–A3) so early instability doesn’t contaminate the noise floor — A4 measures that warm-up explicitly.

- Re-tare before each test block; record whether a re-tare was needed (drift indicator).

- Use actual (weighed) bag masses, never nominal, in analysis. Stacked loads carry ±0.1 g per bag.